## Slide 1
How Much Can Bad Display Calibration Cost?
Calvin Sun1 and Wei-Chung Cheng2
1Johns Hopkins University, Baltimore, MD; 2DIDSR/OSEL/CDRH, U.S. Food and Drug Administration, Silver Spring, MD
The first whole-slide imaging (WSI) device 510(k) de novo was granted on April 12, 2017 based on the Technical Performance Assessment guidance (TPA). Although the importance of color management of the display is stressed in the TPA guidance, many Sponsors are unaware of its significance and consider the display an independent, swappable component. In this poster, we demonstrate the potential errors introduced by mishandling of the color management system.
REGULATORY RELEVANCE
FINDINGS
METHOD
2017 OSEL Annual Student Science Poster Exhibit, August 2, 2017
Fig. 1: Color Management Components.
Clinical Area: Medical Imaging
INTRODUCTION
Unlike grayscale displays, calibration of color medical displays is tightly coupled with the computer environment, which needs to be properly addressed in 510(k) submissions for WSI devices.
CONCLUSION
Experiments demonstrating impacts of mishandling color management for medical displays
ABSTRACT
Experiments were conducted to examine the impact of mishandling color profile, color manager, and graphics card driver for color medical displays. Three color profiles, three color managers, and two graphics card settings were tested. A novel test method using a reconfigurable circuit board was designed and used to measure the output digital pixel values. The experiment results show the entanglement between the color management components from different vendors.
The color management components within the computer environment are depicted in Fig. 1. The basic idea of color management is to compensate for the uncalibrated display errors by modifying the RGB values received from the applications before they are sent to the display. The color profile contains the color transformation for a specific display. The color profile is generated by a color calibration tool. The color manager controls which color profile to use. The graphics card driver executes the color transformation and frequently offers extra features to adjust the display colors. The Pixel Meter is a circuit board that monitors the digital RGB values on the DVI display interface.
What does color calibration do?
Most color calibration tools adjust the gamma curve, gray tracking, and maximum luminance, but not color tracking, minimum luminance, etc.
FUNDING SOURCE
This study was supported by Critical Path Initiative and ORISE.
Calibration Tool | Gamma | Gray Tracking | Max Luminance
Windows | Yes | Yes | No
ColorMunki | Yes | Yes | Yes
Spyder | Yes | Yes | Yes
Chart: Input RGB Value
| 0 | 15 | 30 | 45 | 60 | 75 | 90 | 105 | 120 | 135 | 150 | 165 | 180 | 195 | 210 | 225 | 240 | 255
Windows | 0 | 31 | 52 | 71 | 87 | 103 | 118 | 132 | 146 | 159 | 172 | 185 | 197 | 209 | 221 | 232 | 244 | 255
ColorMunki | 0 | 18 | 34 | 45 | 58 | 69 | 81 | 95 | 109 | 123 | 137 | 152 | 167 | 182 | 198 | 212 | 226 | 241
Spyder | 0 | 45 | 79 | 105 | 125 | 140 | 151 | 160 | 168 | 176 | 183 | 192 | 200 | 209 | 218 | 226 | 232 | 235
Installed Color Manager | ICC profile is decided by
Windows only | latest setting by Windows
Windows and ColorMunki | latest setting by Windows or ColorMunki
Windows and Spyder | latest setting by Spyder
Windows, ColorMunki, and Spyder | latest setting by Spyder
Installed Color Manager | How to choose ICC | Acceptable ICC profiles generated by
Windows | Use “Color Management” | Windows, ColorMunki, Spyder
ColorMunki | Not available | ColorMunki (only the latest calibration)
Spyder4 | Use “ProfileChooser” | Spyder
Are calibrated displays all the same?
No. The calibration results vary greatly among the three calibration tools as shown in the plot.
Which color manager should I use?
You can choose any one but choose only one. Existence of multiple color managers creates confusion for not only the user but also the applications. For example:
How do I use an existing ICC color profile?
Depends on the color managers:
Should I use graphics card software to adjust my display color?
Not recommended because it conflicts with the color profile (e.g., color profile overrides the Gamma setting but not the Digital Vibrance setting).
How much does bad calibration cost?
A 3MP medical display costs $9,500 (Barco MDCC6430). A comparable non-medical display costs $2,000 (NEC EA304WMI-BK-SV). The dE00 between them is 6.12. If the $7,500 difference is paid for the color difference, then the cost is 7500/6.12 = $1,225 per dE00. Thus the worst calibration in our experiments (10.94 dE00) is worse than no calibration and can cost as much as $13,400.
Color
Profile
Graphics Card
Driver
Color
Manager
Computer
Environment
Display
Pixel Meter
Color
Calibration
Tool
Color Profile: generated by Windows, Spyder, or ColorMunki
Color Manager: color profile chosen by Windows, Spyder, or ColorMunki
Graphics Card Software (nVidia Control Panel)
Gamma = 1.0, 1.2, 1.4, or 1.6
Digital Vibrance = 50% or 100%
RESULTS
Color Profile | D.V.=50% | D.V.=100%
None | 0.00 | 7.45
Windows | 7.48 | 10.94
Spyder | 3.81 | 9.90
ColorMunki | 4.35 | 7.96
Use CIEDE2000 (dE00) to calculate the color differences of the 24 color patches
The average of 24 dE00 values are shown in the following table.
24 color patches